EXSI  Details.

ExSI commands are classified in to different categories based on their functionality, the following are the categories

- General Scanner Control

- Exam Control

- Task Control for batch scanning

- Scanner hardware control

- Image data flow control

- Error Handling

- Real time commands

  - General Scanner Control

# ConnectToScanner

The External Scanner Interface shall provide the ability to connect a client to the scanner, provided that the client machine is pre-authenticated by the scanner. The External Scanner Interface shall retain the authenticated clients across scanner software reloads

| Command Name | ConnectToScanner |
| --- | --- |
| Params | product=string |
| Returns | ConnectToScanner=ok¦fail |
| Precondition | Can execute immediately Can execute without exam Can execute during TPS reset |
| Synced | Yes |
| SVAT Command | None |
| Description | This command will be used by client to connect to a specific scanner |
| How it works | Scanner side maintains configuration file that has list of authorized products. Using this command client sends its product name. EXSI server verifies if this product is authorized to connect. If yes it returns ok otherwise it returns fail. |

# GetScannerState

The External Scanner Interface shall provide the ability to retrieve the current scanner status at any given instant

| Command Name | GetScannerState |
| --- | --- |
| Params | None |
| Returns | GetScannerState=ok¦fail scanner=idle¦prescanning¦refscanning¦refprepped¦prepped¦scanning |
| Precondition | None |
| Synced | Yes |
| SVAT Command | None |
| Description | To get the scanner current state |
| How it works | Exsi Server saves the scanner state every time there is a scanner event received from SCAN. GetScanner state command returns this state value to client. Sequence diagram is shown below |

# SetScannerMode

The External Scanner Interface shall provide the ability to change the scanner mode from “clinical” to “research”. However, in order for this command to succeed the research option key should be present, and the scanner should be configured to provide mode change access.

| Command Name | SetScannerMode |
| --- | --- |
| Params | Research |
| Returns | SetScannerMode=ok¦fail |
| Precondition | canExecuteTpsNotReady=false canExecuteWithoutExam=false canExecuteIfPrepped=false canExecuteDuringScanning=false |
| Synced | Yes |
| SVAT Command | PRESC_MODE |
| Description | This command is used to change the scanner mode from clinical to research This command will be used by the user during an exam to move to research mode. |
| How it works | This command looks for the parameter passed. Depending on the parameter passed it calls PRESC_MODE SVAT command with value mapped below as a parameter. Research =1 |

# LockUI

The External Scanner Interface shall provide the ability to prevent users from interacting with the scan desktop on MR operator’s console UI

| Command Name | LockUI |
| --- | --- |
| Params | On ¦ Off |
| Returns | LockUI=ok |
| Precondition | canExecuteTpsNotReady=false |
| Synced | Yes |
| SVAT Command | LOCK_UI, UNLOCK_UI |
| Description | This command provides the capability to put glass pane on the MR host PC UI when a remote operation is in progress to ensure that no UI interaction is allowed. It is possible to view UI & observe all operations. It is also possible to unlock either by using another command with “off” option or by clicking on “unlock UI” button available on console. If system crashes then complete system shuts down closing all of the UI. If external system crashes then user needs to unlock from console. |
| How it works | This command is wrapper to the LOCK_UI & UNLOCK_UI commands in SVAT. If the option is “On”, LOCK_UI command is sent. For “Off” option, UNLOCK_UI command is sent. Default return value from SVAT will be sent directly to client. |

  - Exam Control

# LoadProtocol

The External Scanner Interface shall provide the ability to load any protocols, already installed in the scanner, to the task list. The ability to load the entire protocol, or selected tasks from that protocol shall be given as an option

| Command Name | LoadProtocol |
| --- | --- |
| Params | ge¦site¦service [dir=dir_name] path=protocol_name/task#,task#¦task#-task# |
| Returns | LoadProtocol=ok¦fail taskKeys=string,string,… Returns error if protocol name doesn’t exist at the given location ( ge or site or service / or in the directory mentioned) |
| Precondition | canExecuteTpsNotReady |
| Synced | Yes |
| SVAT Command | PROTOCOL_DIR, PROTOCOL_MODE, LOADPROTOCOL |
| Description | This command is used to load a specific protocol from ge/site/service/directory. |
| How it works | Logical diagram below |

# EndExam

The External Scanner Interface shall provide the ability to end the current exam

| Command Name | EndExam |
| --- | --- |
| Params | None |
| Returns | EndExam=ok¦fail |
| Precondition | canExecuteDuringScanning=false |
| Synced | Yes |
| SVAT Command | END_EXAM |
| Description | This command is used to End currently running exam |
| How it works | This command is a wrapper to the existing SVAT END_EXAM command. Return value from SVAT will be passed to the client. If command is not successful command would return failure. |

# GetExamInfo

The External Scanner Interface shall provide the ability to retrieve details of the current exam. The details shall include, but not limited to the following items:

Exam number, Patient Name, ID, Date of Birth, Sex, Weight, Operator Name, Exam Description

| Command Name | GetExamInfo |
| --- | --- |
| Params | None |
| Returns | GetExamInfo=ok¦fail with patient/exam info encoded as DICOM tag=value pairs, e.g. 0010,0010=string 0010,0020=string Study id, Patient name, Patient ID, Patient DOB, Sex, Age, Weight, Operator name, Radiologist, Ref physician, description, Additional patient history Eg. GetExamInfo=ok 0020,0010=40 0010,0010=patient^Name 0010,0020=1234 0010,0030=20010923 0010,0040=M 0010,1010=013Y 0010,1030=44.000 0008,1070=operator 0008,1060=radiologist 0008,0090=dr |
| Precondition | canExecuteDuringScanning=false |
| Synced | Yes |
| SVAT Command | None |
| Description | This command is used to get current exam information. |
| How it works | Logical diagram is shown below. |

# DisablePrompts

The External Scanner Interface shall provide the ability to set exam level options for the duration of current exam. These options shall include the following:

a. Disabling all advisory errors and pop-ups (Disabling the modal pop-ups, which require user acknowledgement from the MR operator console shall allow the continuous flow of remote operations)

| Command Name | DisablePrompts |
| --- | --- |
| Params | none ¦ advisory |
| Returns | DisablePrompts=ok¦fail |
| Precondition | default |
| Synced | Yes |
| SVAT Command | SCN_POPUPS |
| Description | This command is used to Disable prompts for All/Grx/Advisory/auto_start. |
| How it works | Suppressed pop-ups are Grx SWIFT popup, auto start and advisory. None of them are related to hazard. Grx SWIFT popup & auto start pop are warning pop-up. There is no impact if they are suppressed. Suppressing advisory pop-up would set the advisory values automatically. Logical diagram is shown below |

# GetTaskList

The External Scanner Interface shall provide the ability to retrieve the list of tasks currently available in the Task List, preserving the order in which they are listed, along with a unique key for each task.

| Command Name | GetTaskList |
| --- | --- |
| Params | none |
| Returns | GetTaskList=ok¦fail taskKeys=string,string,… |
| Precondition | canExecuteTpsNotReady=false, canExecuteDuringScanning=false, canExecuteIfPrepped=false |
| Synced | Yes |
| SVAT Command | None |
| Description | This command is used to Get all task displayed in the task list. |
| How it works | Logical diagram is shown below |

# SelectTask

The External Scanner Interface shall provide the ability to select a task by providing the unique key of that task

| Command Name | SelectTask |
| --- | --- |
| Params | taskKey=string |
| Returns | SelectTask=ok¦fail Returns following errors: -no task key given by user -incorrect task key |
| Precondition | Default |
| Synced | Yes |
| SVAT Command | None |
| Description | This command is used to select the specific task entered as task key. |
| How it works | Logical diagram is shown below |

# ActivateTask

The External Scanner Interface shall provide the ability to select a task by providing the unique key of that task

| Command Name | ActivateTask |
| --- | --- |
| Params | taskKey=string |
| Returns | ActivateTask=ok¦fail returns failures: -If task cannot be selected for given task key. -the current task state does not permit activate operations if task is in scanned state, View edit fails, save fails, download fails |
| Precondition | Cannot execute when prepped, TPS not ready, during scanning |
| Synced | No |
| SVAT Command | VIEW_EDIT, ACCEPT_RX, SCANOPS_EVENT |
| Description | This command moves a task to the ACT state depending on its current state. |
| How it works | This command is a meta command executes the above mentioned 3 SVAT command depending on its current state. If task is in NEW state = > VIEW_EDIT->ACCEPT_RX->SCANOPS_EVENT If task is in INRX state => ACCEPT_RX->SCANOPS_EVENT If task is in RXD state => SCANOPS_EVENT If task is in ACT state => Failure message |

  - Task Control for batch scanning

# setRxParams

The External Scanner Interface shall provide the ability to set or modify prescription parameters, including the slice collection for the given task key

The External Scanner Interface shall provide the ability to set a coil using task key

| Command Name | setRxParams |
| --- | --- |
| Params | [taskKey=string] [coil=string] [TR=float] [TE=float] [TI=float] [ETL=int] [NUMECHO=int] [NEX=float] [FLIP=int] [BW=float] [XRES=int] [YRES=int] |
| Returns | SetRxParams=ok¦fail |
| Precondition | Cannot execute if TPS is not ready |
| Synced | Yes |
| SVAT Command | VIEW_EDIT to setUp the task. |
| Description | Sets the general prescription parameters as described above. (a) If a task Key is provided, the task is selected in the task list and view edited if allowed to be modified. (b) If no task key is provided, the selected task is view edited if allowed. The desired parameters are set if the view edit is successful. The params value setting continues for other params even if one/more of the requested params setting fails. |
| How it works | Diagram is shown below |

# setRxGeometry

The External Scanner Interface shall provide the ability to set or modify prescription parameters, including the slice collection for the given task key

| Command Name | setRxGeometry |
| --- | --- |
| Params | [taskKey=string] [coil=string] [plane=axial/sagittal/coronal/oblique] [startloc=int¦r,a,s] [endloc=int¦r,a,s] [slices=int] [spacing=int] [thick=int] [fov=int] [freqdir=swap/unswap/AP/RL/SI] where startloc and endloc should be in r,a,s only for oblique scan plane |
| Returns | SetRxGeometry=ok¦fail |
| Precondition | Cannot execute if TPS is not ready |
| Synced | Yes |
| SVAT Command | Plane: SCAN_PLANE Coil: SET_COIL Startloc: START_LOCS, Endloc: END_LOCS Slices: NO_2D_SLICES Spacing: SLICE_SPACING Thick: SLICE_THICKNESS Fov: FOV Freqdir: FREQ_DIR |
| Description | Sets the parameters related to geometry like plane, coil, start location, end location, slices, spacing, thickness, fov and freqdir. (a) If a task Key is provided, the task is selected in the task list and view edited if allowed to be modified. (b) If no task key is provided, the selected task is view edited if allowed. The desired parameters are set if the view edit is successful.. |
| How it works | ExSI server maps the input params to the legacy SVAT commands. The param value setting continues for other params even if one/more of the requested params setting fails. Here is the order in which parameters are set 1. SCAN_PLANE 2. SET_COIL 3. START_LOCS 4. END_LOCS 5. NO_2D_SLICES 6. SLICE_SPACING 7. SLICE_THICKNESS 8. FOV 9. FREQ_DIR |

# setRxGeometry3p

The External Scanner Interface shall provide the ability to set or modify prescription parameters, including the slice collection for the given task key

| Command Name | setRxGeometry3p |
| --- | --- |
| Params | [taskKey=string] [coil=string] [centerloc=r,a,s] [spacing=r,a,s] [thick=int] [fov=int] [freqdir=swap/unswap] |
| Returns | SetRxGeometry3p=ok¦fail |
| Precondition | Cannot execute if TPS is not ready |
| Synced | Yes |
| SVAT Command | Coil: SET_COIL Centerloc: SET_SL_CENTER Spacing: SET_SL_SPACING Thick: SLICE_THICKNESS Freqdir: FREQ_DIR Plane: SCAN_PLANE (always send 3-plane) |
| Description | Sets the parameters related to geometry for 3 plane like plane(always send 3-plane), coil, center location, spacing, thickness, fov and freqdir. (a) If a task Key is provided, the task is selected in the task list and view edited if allowed to be modified. (b) If no task key is provided, the selected task is view edited if allowed. The desired parameters are set if the view edit is successful. |
| How it works | ExSI server maps the input params to the legacy SVAT commands. The param value setting continues for other params even if one/more of the requested params setting fails. |

# prescan

The External Scanner Interface shall provide the ability to perform

- auto prescan b. manual prescan c. skip prescan d. set prescan values while showing/hiding  prescan UI

| Command Name | prescan |
| --- | --- |
| Params | auto¦manual=start/done ¦ skip ¦ values=hide/show [r1=int] [r2=int] [tg=int] [cf=int] [x=int] [y=int] [z=int] |
| Returns | Prescan=ok¦fail |
| Precondition | Cannot execute if TPS is not ready, Cannot execute if scanner is in prepped state Cannot execute during acquisition scanning and ref scanning, but can execute during prescanning. |
| Synced | Auto Prescan : Synced Manual Prescan start/stop :Synced Skip Prescan, setting of prescan/shim value : Not Synced |
| SVAT Command | auto :APS_EVENT manual value=start/stop :START_LOOP_EVENT, STOP_LOOP_EVENT skip :PSC_MPS_EXIT_MPS values=hide [r1=] [r2=] [tg=int] [cf=int] [x=int] [y=int] [z=int] :PSC_UPDATE_VAL, SC_SET_SHIM, PSC_MPS_EXITMPS values=show [r1=] [r2=] [tg=int] [cf=int] [x=int] [y=int] [z=int] :START_LOOP_EVENT, PSC_UPDATE_VAL, PSC_SET_SHIM, STOP_LOOP_EVENT |
| Description | Prescan can be performed using any of the below mentioned options. 1) Manual Prescan without any prescan value modification: Prescan manual=start/done 2) Auto Prescan: Prescan auto 3) Skip Prescan: Prescan skip 4) Set prescan values while showing the manual prescan UI as done from user interface: Prescan values=show [r1=int] … 5) Set prescan values without showing manual prescan UI: Prescan values=hide [r1=int] … |
| How it works | Details in the figure below |

# GetPrescanValues

The External Scanner Interface shall provide the ability retrieve prescan values of the current task after performing a Prescan operation

| Command Name | GetPrescanValues |
| --- | --- |
| Params | None |
| Returns | GetPrescanValues=ok¦fail stat=[ok¦fail¦prescanning] r1=int r2=int tg=int cf=int |
| Precondition | Cannot execute if TPS is not ready |
| Synced | Yes |
| SVAT Command | PSC_QUERY_STATUS |
| Description | Returns the status of prescan e.g. prescanning, last prescan was successful or not. Returns the prescan values (r1, r2, tg, cf)of the last prescan if it was successful. |
| How it works | Please refer the diagram for more details |

# ReferenceScan

The External Scanner Interface shall provide the ability to perform “reference scan” for operations required prior to beginning acquisition

| Command Name | ReferenceScan |
| --- | --- |
| Params | None |
| Returns | ReferenceScan=ok¦fail |
| Precondition | Cannot execute if TPS is not ready, Cannot execute if scanner is in prepped state Cannot execute if scanning is in progress |
| Synced | No |
| SVAT Command | REF_EVENT |
| Description | Performs reference scan. Performs download/full(or smart) prescan if the task has not been already downloaded/prescanned. Ref-scans the selected task in the task list UI. |
| How it works | ExSI server maps the command to the legacy SVAT command for reference scan. |

# Prep

The External Scanner Interface shall provide the ability to perform “scan prep” for operations required prior to beginning acquisition.

| Command Name | Prep |
| --- | --- |
| Params | None |
| Returns | Prep=ok¦fail |
| Precondition | Cannot execute if TPS is not ready, Cannot execute if scanner is in prepped state Cannot execute if scanning is in progress |
| Synced | No |
| SVAT Command | PREP_SCANNER |
| Description | Performs prep. Performs download/full(or smart) prescan if the task has not been already downloaded/prescanned. Preps the selected task in the task list UI. |
| How it works | ExSI server maps the command to the legacy SVAT command for prep. |

# Scan

The External Scanner Interface shall provide the ability to start scanning any prescribed and saved task of the current exam by providing the unique task Key associated with that task.

The External Scanner Interface shall provide the ability to rescan an already scanned task of the current exam. The requested task should be the currently active task

| Command Name | Scan |
| --- | --- |
| Params | None |
| Returns | Scan=ok¦fail |
| Precondition | Cannot execute if TPS is not ready, Cannot execute if scanning is in progress |
| Synced | No |
| SVAT Command | SCAN_EVENT |
| Description | Performs scan. Performs download/full(or smart) prescan/prep if the task has not been already downloaded/prescanned/prepped. Rescans an already scanned selected task. If the scanned task is active, rescans and generates images in the same task. If the scanned task is not active, duplicates the task, scans and generates images in the new task. |
| How it works | ExSI server maps the command to the legacy SVAT command for scan. |

# StopScan

The External Scanner Interface shall provide the ability to immediately stop scanning (including pre-scanning) of the current task. This command shall be equivalent to the Stop Scan Button on the hardware panel (hard keys). This command shall have no effect if a scanning task is unavailable at the time of receiving the command.

| Command Name | StopScan |
| --- | --- |
| Params | None |
| Returns | StopScan=ok¦fail |
| Precondition | Cannot execute if TPS is not ready, Cannot execute if scanning is in progress |
| Synced | No |
| SVAT Command | STOP_SCAN |
| Description | Stops all kinds of scanning like prescanning, ref scanning, scanning. |
| How it works | ExSI server maps the command to the legacy SVAT command for stop scan. |

# setCvs

The External Scanner Interface shall provide the ability to modify control variables (CV) for the currently downloaded task. All CVs shall be allowed to be modified in research mode. Only pre-defined CVs for pre-defined psds shall be allowed to be modified in clinical mode

| Command Name | setCvs |
| --- | --- |
| Params | name1=string value name2=string value … |
| Returns | setCvs=ok¦fail taskkey = string name1=string value name2=string... |
| Precondition | Cannot execute if scanner is in prepped state. Cannot execute if scanning is in progress. Cannot execute if TPS not ready |
| Synced | Yes |
| SVAT Command | None |
| Description | All CVs shall be allowed to be modified in research mode. Only pre-defined CVs for pre-defined psds shall be allowed to be modified in clinical mode. |
| How it works | A CV access controller class shall contain the list of CVs and PSDs allowed to be modified. The data in the class could be embedded in the code or obtained from an external file. The CV setting continues for other CVs even if one/more of the requested CV setting fails. |

# getCvs

The External Scanner Interface shall provide the ability to retrieve control variable (CV) values for a currently downloaded task

| Command Name | getCvs |
| --- | --- |
| Params | name1 name2 … |
| Returns | getcvs=ok¦fail name1=string value... name2=string value |
| Precondition | Cannot execute if TPS not ready |
| Synced | Yes |
| SVAT Command | None |
| Description | Provides the CV values for the currently active task. |
| How it works | Sends a geocomm request to ScannerService and receives the CV values in the response. |

  - Scanner hardware control.

# GetConnectedCoils

The External Scanner Interface shall provide the ability to retrieve a list of all connected coils.

| Command Name | GetConnectedCoils |
| --- | --- |
| Params | None |
| Returns | GetConnectedCoils=ok¦fail coils=string[,string,...] |
| Precondition | Can execute without Exam Cannot execute when TPS not ready |
| Synced | Yes |
| SVAT Command | CONNECTED_COILS |
| Description | This command returns all connected coils. |
| How it works | This command sends SVAT command CONNECTED_COILS to get list of all connected coils. |

# PatientTable

The External Scanner Interface shall provide the ability to advance the cradle in order to position the scan location at isocenter

The External Scanner Interface shall provide the ability to set landmark

The External Scanner Interface shall provide the ability to turn on/off any cradle movements.

| Command Name | PatientTable |
| --- | --- |
| Params | advanceToScan autoLandmark=on/off emulation=on/off freeze=on/off landmark=int |
| Returns | PatientTable=ok¦fail |
| Precondition | Can execute without Exam Cannot execute when TPS not ready Cannot execute during scanning |
| Synced | No |
| SVAT Command | IPG_ADVANCE_TOSCAN, LANDMARK, AUTOLANDMARK, IPG_TABLE_EMULATION FREEZE_TABLE |
| Description | This command provides ability to turn on/off any cradle movements, automatically move cradle to desired position |
| How it works | This command sends SVAT command IPG_ADVANCE_TOSCAN to move the cradle to desired position. LANDMARK to set the landmark. FREEZE_TABLE to turn on/off the table movement. |

  - Image data flow control.

# GenerateImageTypes

The External Scanner Interface shall provide the ability to enable/disable generating the Magnitude, Phase, Imaginary & Real image types in any combination.

| Command Name | GenerateImageTypes |
| --- | --- |
| Params | all ¦ [magnitude] [phase] [real] [imaginary] |
| Returns | GenerateImageTypes=ok¦fail |
| Precondition | Cannot execute if scanner is in prepped state. Cannot execute if scanning is in progress. Cannot execute if TPS not read |
| Synced | Yes |
| SVAT Command | None |
| Description | Generates all or few of different image types that is magnitude, phase, real and imaginary in any combination. The command works on currently active task. |
| How it works | Sets the CV rhrcctrl “bitMask to indicate types of images to be generated” . Here are the bit values to be set: Magnitude: 0x1 Phase: 0x2 Real: 0x4 Imaginary: 0x8 Please refer setCv command to find details of CV setting mechanism. |

# installImageTypes

The External Scanner Interface shall provide the ability to enable/disable saving non-realtime images to the database

| Command Name | installImageTypes |
| --- | --- |
| Params | all ¦ none ¦ [magnitude] [phase] [real] [imaginary] |
| Returns | InstallImageTypes=ok¦fail |
| Precondition | Cannot execute if scanner is in prepped state. Cannot execute if scanning is in progress. Cannot execute if TPS not ready |
| Synced | Yes |
| SVAT Command | None |
| Description | Installs all or few of different image types (that is magnitude, phase, real and imaginary in any combination) if the type is being generated. It works on active task. Therefore the command should be issued post download. Settings are applicable on the images generated from next scan. Rescan of the same task would retain these settings. |
| How it works | Creates bit mask from the input params indicating the following: Magnitude: 0x1, Phase: 0x2, Real: 0x4, Imaginary: 0x8 The bitmask is sent from EXSI server to ScannerService. In scn process a packet containing exam number, series number, scan number and bit mask is created and sent to TIR to block image installation as requested. |

# GenerateIntermediateImages

The External Scanner Interface shall provide the ability to enable/disable generating images from each channel of the prescribed coil configuration. This feature will be effective only for multi-channel coils. If enabled, an image from each enabled “image type” shall be provided from each channel. In addition, an image for each enabled “image type” shall be provided from all channels combined.

| Command Name | GenerateIntermediateImages |
| --- | --- |
| Params | on¦off |
| Returns | GenerateIntermediateImages=ok¦fail |
| Precondition | Cannot execute if scanner is in prepped state. Cannot execute if scanning is in progress. Cannot execute if TPS not ready |
| Synced | Yes |
| SVAT Command | None |
| Description | Generates images from all channels of the selected coil along with combined images. The command works on currently active task |
| How it works | sets CV "saveinter“ to value 1(0) as per the user input of on(off) Please refer setCv command to find details of CV setting mechanism. |

  - Error Handling

# NotifyEvent

The External Scanner Interface shall provide the ability to register and deregister for scanner state changes. When registered, the client shall be notified when a scanner state change occurs.

Upon a successful acquisition, the External Scanner Interface shall send the “Acquisition Complete” event to the client, which has registered for this event.

When batch image receive components submit the first image(of a task) to the database, the External Scanner Interface shall send the “First Image Available” event to the client along with total number of images expected, which  has registered for this event

When batch image receive components submit the last image (of a task) to the database,, the External Scanner Interface shall send the “All Images Available” event to the client, which  has registered for this event

The External Scanner Interface shall return an error response to synchronous commands which did not execute successfully. The error response shall contain pre-defined error codes and optional translated character string (in the system default character set) which can optionally be presented to the operator by the external client.

The External Scanner Interface shall broadcast error messages when known exceptions and error conditions occur within the scanner. A mechanism shall be provided for the clients to register and get notified on these error conditions.

| Command Name | NotifyEvent |
| --- | --- |
| Params | NotifyEvent [all=on¦off] ¦ [acquisition=on¦off] [error=on¦off] [exam=on¦off] [rxError=on¦off] [recon=on¦off] [scanmode=on¦off] [scanner=on¦off] [tps=on¦off] [firstImage=on¦off] |
| Returns | notifyEvent = ok¦fail [all=on¦off] ¦ [acquisition=on¦off] [error=on¦off] [exam=on¦off] [rxError=on¦off] [recon=on¦off] [scanmode=on¦off] [scanner=on¦off] [tps=on¦off] [firstImage=on¦off] |
| Precondition | Can be executed without exam Can be executed immediate Can be executed during TPS Reset |
| Synced | Yes |
| SVAT Command | None |
| Description | Notifies the client for the registered events. The client can register for all or some of these events: Scanner state change (idle/prescanning/prepped/refprepped/scanning/refscanning) Scanner mode (clinical/research) change TPS state (resetting/reset_failed) Exam state (started/ended) Acquisition done First Image available(submitted to database) with total # of expected images All images available(last image submitted to database, i.e. recon done) with actual # images processed and sent to database. RxError to indicate the errors that are shown on the UI during prescription like parameter not filled etc. Error to indicate predefined errors for which client has registered using ermes or some other mechanism. For e.g. error to indicate that an invalid CV set has been attempted etc. The details to be included in the Error Handling requirements/design.images. The command works on currently active task |
| How it works | Please refer to images below. |

  - Real time commands

# SetRealtimeOptions

The External Scanner Interface shall provide the ability to start realtime scanning with or without displaying real-time acquisition screen.

| Command Name | SetRealtimeOptions |
| --- | --- |
| Params | EXSI_CONTROLLER = EXSI_HOST ¦EXSI_REMOTE |
| Returns | SetRealtimeOptions=ok ¦fail |
| Synced | Yes |
| Workflow | This command will be used by the user before starting real time scan. EXSI_CONTROLLER = EXSI_HOST This will enable user interface when real time scanning starts EXSI_CONTROLLER = EXSI_REMOTE This will hide user interface when real time scanning starts |
| Description | This command will enable the system to start real time scanning with or without real time user interface. |
| How it works | EXSI sends GEOCOM request to SCN, which would check if system is idle and there is active task available. If not it returns error. If yes then it sets the flag. This flag is used to hide/show RTIA UI when scan is requested. |

# PauseScan

The External Scanner Interface shall provide the ability to pause the current real time acquisition after completing the current pass.

| Command Name | PauseScan |
| --- | --- |
| Params | none |
| Returns | PauseScan=ok¦fail |
| Precondition | Cannot execute if TPS is not ready, Cannot execute without exam Can execute if real time scanning is in progress |
| Synced | Yes |
| SVAT command | PAUSE_EVENT |
| Description | If realtime scanning is in progress it pauses the scan. If scanner is not scanning it returns fail with error “"scanner is not scanning“ If scanner is already paused, it returns fail with error "scanner is already paused" If scanner is not realtime, it returns fail with error "series is not realtime series" |
| How it works | It uses SVAT command PAUSE_EVENT |

# ResumeScan

The External Scanner Interface shall provide the ability to resume a paused realtime acquisition.

| Command Name | ResumeScan |
| --- | --- |
| Params | none |
| Returns | ResumeScan=ok¦fail |
| Precondition | Cannot execute if TPS is not ready, Cannot execute without exam Can execute if real time scanning is in progress |
| Synced | Yes |
| SVAT command | RESUME_EVENT |
| Description | If realtime scanning is in paused state it resumes the scan. If scanning is not paused, it returns fail with message “scanner is not paused” If task is not real time, it return fail with message “series is not realtime series” |
| How it works | It uses svat command RESUME_EVENT |

# GenerateIntermediateImages

The External Scanner Interface shall provide the ability to generate real-time images for all channels of coil

| Command Name | GenerateIntermediateImages |
| --- | --- |
| Params | on ¦ off |
| Returns | GenerateIntermediateImages=ok¦fail |
| Precondition | Cannot execute if scanner is in prepped state. Cannot execute if scanning is in progress. Cannot execute if TPS not ready |
| Synced | Yes |
| SVAT command | None |
| Description | This command controls generation of images from all channels of coil during real time scan. |
| How it works | It sets saveinter CV to generate images from all channels |

# GenerateImageTypes

The External Scanner Interface shall provide the ability to enable/disable generating the Magnitude, Phase, Imaginary & Real image types in any combination.

| Command Name | GenerateImageTypes |
| --- | --- |
| Params | all ¦ [magnitude] [phase] [real] [imaginary] |
| Returns | GenerateImageTypes=ok¦fail |
| Precondition | Cannot execute if scanner is in prepped state. Cannot execute if scanning is in progress. Cannot execute if TPS not ready |
| Synced | Yes |
| SVAT command | None |
| Description | Generates all or few of different image types that is magnitude, phase, real and imaginary in any combination. The command works on currently active task. |
| How it works | See the setting of rhrctctrl below. |

# installImageTypes

The External Scanner Interface shall provide the ability to enable/disable saving realtime images to the database.

| Command Name | installImageTypes |
| --- | --- |
| Params | all ¦ none ¦ [magnitude] [phase] [real] [imaginary] |
| Returns | InstallImageTypes=ok¦fail |
| Precondition | Cannot execute if scanner is in prepped state. Cannot execute if scanning is in progress. Cannot execute if TPS not ready |
| Synced | Yes |
| SVAT command | None |
| Description | Installs all or few of different image types (that is magnitude, phase, real and imaginary in any combination) if the type is being generated. It works on active task. Therefore the command should be issued post download. Settings are applicable on the images generated from next scan. Rescan of the same task would retain these settings. |
| How it works | Creates bit mask from the input params indicating the following: Magnitude: 0x1, Phase: 0x2, Real: 0x4, Imaginary: 0x8 The bitmask is sent from EXSI server to ScannerService. In scn process a packet containing exam number, series number, scan number and bit mask is created and sent to RIR to block image installation as requested. |

# RealTimeImages

The External Scanner Interface shall provide the ability to enable/disable sending realtime images to the Auto-view window. Images sent to auto-view shall be shown immediately unless the auto-view is configured to be on the review mode

| Command Name | RealTimeImages |
| --- | --- |
| Params | [autoview=on¦off] |
| Returns | RealTimeImages=ok¦fail |
| Precondition | Cannot execute if scanner is in prepped state. Cannot execute if TPS not ready |
| Synced | Yes |
| SVAT command | None |
| Description | This command controls real time image display in autoview. autoview=on This will display real time images in autoview autoview=off This will stop displaying images in autoview |
| How it works | See the setting of rhrctctrl below. |

# BeginRealtimeUpdate

The External Scanner Interface shall provide commands to begin and end real time parameter changes. The External Scanner Interface shall queue all parameter updates from a “BeginRealtimeUpdate” command until an “EndRealtimeUpdate” command is received. Table

A “Begin Realtime Update” shall discard any prior uncommited changes. “End RealtimeUpdate” command shall provide the option to commit or discard accumulated realtime parameter updates.

| Command Name | BeginRealtimeUpdate |
| --- | --- |
| Params | none |
| Returns | BeginRealtimeUpdate=ok¦fail |
| Synced | Yes |
| SVAT command | None |
| Description | The External Scanner Interface shall queue all parameter updates from a “BeginRealtimeUpdate” command until an “EndRealtimeUpdate” command is received. |
| How it works | This command will create new object for including all real time requests. It will also reset the counter used to track number of requests. |

# EndRealtimeUpdate

The External Scanner Interface shall provide commands to begin and end real time parameter changes. The External Scanner Interface shall queue all parameter updates from a “BeginRealtimeUpdate” command until an “EndRealtimeUpdate” command is received. Table

A “Begin Realtime Update” shall discard any prior uncommited changes. “End RealtimeUpdate” command shall provide the option to commit or discard accumulated realtime parameter updates.

| Command Name | EndRealtimeUpdate |
| --- | --- |
| Params | commit ¦ discard |
| Returns | EndRealtimeUpdate=ok¦fail |
| Synced | Yes |
| Workflow | This command will be used after sending all realtime requests. Commit: This argument will send all parameter updates to real time application for processing. Discard: This argument will not send any parameter update. It cleans up all requests in the queue. |
| Description | The External Scanner Interface shall queue all parameter updates from a “Begin Realtime Update” command until an “End Realtime Update” command is received. |
| How it works | If commit is passed then it sends all parameter requests in the queue to RTPSI for processing. It returns the result after processing. If discard is passed it deletes the object that holds all requests in the queue. Resets the request counter. |

# SetRealtimeParams

SWRS.EXSI.214: The External Scanner Interface shall provide the ability to send the following list of real time parameters to an active realtime series.

1. FOV

2. Slice thickness

3. Frequency Direction (along row/column)

4. Flip Angle

5. Track mode

6. Trip related parameters

7. Discarded Data Acquisition

8. Parameters for auto test framework

These parameters shall be queued in the active series until a separate command is sent to commit

the changes.

| Command Name | SetRealtimeParams |
| --- | --- |
| Params | [centerloc=r,a,s rowvec=r,a,s columnvec=r,a,s] [flip=int] [fov=int] [freqdir=" EXSI_ROW "¦" EXSI_COLUMN "] [thick=int] [track=" EXSI_ON"] [trip_nex=int] [dda=int] [trig=" EXSI_ON "¦" EXSI_OFF "] [tripRef=" EXSI_ON"] [trip_fov1=float] [trip_fov2=float] [trip_fov3=float] [trip_rlocsi=float] [trip_rlocap=float] [trip_rlocrl=float] [rotate=int transpose=int path=string] |
| Returns | SetRealtimeParams=ok ¦fail |
| Synced | Yes |
| Workflow | This command will be used for sending all real time requests. [centerloc=r,a,s rowvec=r,a,s columnvec=r,a,s] All the three arguments must be provided together. First one is center of location to be acquired. Second and third argument will give row and column vector respectively. [fov=int] This is to set the fov [thick=int] This is to set the thickness [freqdir=" EXSI_ROW "¦" EXSI_COLUMN "] This is to set the frequency direction [track=on] set to 1 to request a tracking [dda=int] to control the first phase imaging dda after switching back from tracking [trip_nex=int] nex value in trip [trig=on/off] to switch external trigger on/off during realtime imaging [trip_fov1=float] the first pass FOV of the tracking PSD [trip_fov2=float] the second pass FOV of the tracking PSD [trip_fov1=float] the third/fourth pass FOV of the tracking PSD [trip_rlocap] the AP center of the tracking PSD [trip_rlocsi] the SI center of the tracking PSD [trip_rlocrl] the RL center of the tracking PSD [tripRef=on] set to 1 to request a reference scan in realtime EPI [rotate=int transpose=int path=string] rotate angle. Transpose and golden data path for geometry testing |
| Description | This command shall add real time parameter request into queue. |
| How it works | If commit is passed then it sends all parameter requests in the queue to RTPSI for processing. It returns the result after processing. If discard is passed it deletes the object that holds all requests in the queue. Resets the request counter |